## Slide 1: SCD4X
Breaking the size barrier in optical CO2 sensing

## Slide 2: SCD4x
Up to
10x smaller
than other optical CO2 sensors
Photoacoustic NDIR
Selective, high-accuracy measurement
Cost Effective and reflow solderable
The world’s smallest optical CO2 sensor line combines cost-effectiveness with accuracy and versatility. Three variants available to cover a range of performance requirements and operation modes.
Breaking the size barrier in optical CO2 sensing
SCD30
SCD4x
Robust
against dust and liquid water

## Slide 3: SCD4x CO2 Sensor Family
Product Introduction

## Slide 4: SCD4x: a truly disruptive CO2 sensor
  SCD4x Features
Up to 10 times smaller than other optical CO2 sensors!
World’s smallest optical CO2 sensor: 10.1 x 10.1 x 6.5 mm
© Copyright Sensirion, Switzerland
4

## Slide 5: SCD4x: a truly disruptive CO2 sensor
  SCD4x Features
Automated, low cost and reliable assembly
World’s smallest optical CO2 sensor
SMD solderable, tape & reel package
© Copyright Sensirion, Switzerland
5

## Slide 6: SCD4x: a truly disruptive CO2 sensor
  SCD4x Features
Highly selective sensing, factory calibrated & on-board compensated
World’s smallest optical CO2 sensor
SMD solderable, tape & reel package
Great sensing performance
Output range of up to: 0 ppm – 40’000 ppm
Specified range of up to: 400 – 5000 ppm
Specified accuracy of up to: ±(30ppm + 3%)
Fully calibrated and linearized output
Comprehensive on-board temperature & humidity compensation
Intrinsic compensation of CO2 output with altitude/pressure input possible
© Copyright Sensirion, Switzerland
6

## Slide 7: SCD4x: a truly disruptive CO2 sensor
  SCD4x Features
Robust against mechanical stress, dust and liquid water
World’s smallest optical CO2 sensor
SMD solderable, tape & reel package
Great sensing performance
Unmatched sensor robustness
© Copyright Sensirion, Switzerland
7
Encapsulated sensor design, no exposed electronic components
Successfully passed qualification tests according to IP57 (SCD4x sensor component only)

## Slide 8: SCD4x: a truly disruptive CO2 sensor
  SCD4x Features
Highest sensor versatility
World’s smallest optical CO2 sensor
SMD solderable, tape & reel package
Great sensing performance
Unmatched sensor robustness
Large supply voltage range: 2.4 V – 5.5 V
Low power mode and single shot mode available
Includes RH and T output signal
© Copyright Sensirion, Switzerland
8

## Slide 9: SCD41: Improved Accuracy, Extended Range & Single-Shot Mode
  SCD41 Features
Extended Specified Range
Small CO2 sensor size
SMD solderable, tape & reel package
On-chip RH and T compensation
Large supply voltage range
Critical components by Sensirion → ensured sensor availability
CO2 accuracy: ± (50 ppm + 2.5 % m.v.) @ 400–1’000 ppm ± (50 ppm + 3 % m.v.) @ 1001–2’000 ppm ± (40 ppm + 5 % m.v.)
@ 2001–5’000 ppm
Automatic Self-Calibration (ASC) eliminates need for manual field calibration
Features single-shot mode operation
© Copyright Sensirion, Switzerland
9
Single-Shot Operation Mode

## Slide 10: SCD43: Comprehensive Building Standard Compatibility
  Key Features
Compatible with: ASHRAE 62.1-2022 Addendum d
California Title 24
RESET®
WELL Building Standard™
Small CO2 sensor size
SMD solderable, tape & reel package
On-chip RH and T compensation
Large supply voltage range
Critical components by Sensirion → ensured sensor availability
“Best-in-class” CO2 accuracy: ± (30 ppm + 3 % m.v.) @ 400 – 5’000 ppm
Automatic Self-Calibration (ASC) eliminates need for manual field calibration
Features single-shot mode operation
© Copyright Sensirion, Switzerland
10

## Slide 11: Overview: SCD4x Product Variants
SCD40
SCD41
 | Product variant for cost sensitive markets | Product variant for improved accuracy, extended range and/or single shot modes | Product variant for comprehensive building standard compatibility
Accuracy | ± (50 ppm + 5 % of reading) | ± (50 ppm + 2.5 % of reading)* ± (50 ppm + 3 % of reading)** ± (40 ppm + 5 % of reading)*** | ± (30 ppm + 3 % of reading)
Specified output range | 400 – 2’000 ppm | *400 – 1’000 ppm **1001 – 2’000 ppm ***2001 – 5’000 ppm | 400 – 5’000 ppm
Output range | 0 – 40’000 ppm |  | 
Form / footprint / interface | 10.1 x 10.1 x 6.5 mm (nominal), digital I2C interface |  | 
Operation modes | Periodic (5s) Low-Power Periodic (30s) | Periodic (5s) Low Power Periodic (30s) Single-shot | 
© Copyright Sensirion, Switzerland
11
SCD43

## Slide 12: Photoacoustic NDIR
The innovative technology inside SCD4x

## Slide 13: SCD4x: Photoacoustic NDIR Technology
Narrowband IR emission targeting CO2 absorption band
Selective absorption of IR light by CO2 molecules
De-excitation leading to pressure increase in the measurement cell
Periodic modulation of the source to obtain the photoacoustic signal
Low CO2 concentration
High CO2 concentration
© Copyright Sensirion, Switzerland
13

## Slide 14: PASens® technology – competitive edge through innovation
Component protection through metal cap  Highest robustness against dust, condensation and mechanical impact
Sensirion MEMS IR emitter  Good long-term stability
Sensirion humidity and temperature sensor  Superior compensation across a wide RH and T range
Sensirion signal processor
 Advanced signal processing for high accuracy and low noise level
© Copyright Sensirion, Switzerland
14

## Slide 15: Yes, this is all Real CO2
eCO2 is a Derived Estimation
eCO2 measurement is derived from the reactions of all substances in the air, measured with a MOX sensor
Key assumption is that Hydrogen, as VOC produced by humans is proportional to their exhaled CO2
Strong VOC events create false positives. PCC*<0 (1)  flipping a coin would have produced a better result for the given test setting
eCO2 = f( + + + + + +… )
CO2 is a Direct Measurement
Spectroscopy NDIR measures how much infrared light is absorbed by CO2 molecules
Photoacoustic PA NDIR measures pressure waves generated by CO2 molecules which absorbed infrared light
Thermal conductivity TC measures thermal conductivity in air due to changes in CO2 concentration
There is no alternative to a dedicated CO2 sensor (1)
*PCC = Pearson Correlation Coefficient
Notes: I1
sources:
Performance assessment of low-cost environmental monitors and single sensors under variable indoor air quality and thermal conditions https://www.sciencedirect.com/science/article/pii/S0360132320307836

## Slide 16: Making The Difference With CO2 Sensing
SCD4x Use-Cases

## Slide 17: Indoor CO2 Detection Use Cases
How CO2 sensing can bring value to our customers
Healthy Indoor Environments High levels of CO2 can lead to a reduced sense of well-being as well as lower cognitive performance.
Viral Index Higher CO2 concentrations have been shown to be correlated with a higher risk of infection.
Presence Detection Anonymous and fast detection of CO2 concentration can serve as a complement to traditional presence detectors.
Indoor Environment Changes Rapidly changing temperatures and decreasing CO2 levels can indicate an open window or door.
Energy Efficiency
  Monitoring CO2 levels enables optimization of ventilation and lighting to reduce energy consumption.
© Copyright Sensirion, Switzerland
17
Notes: B1

## Slide 18: CO2 is a Major Differentiator
Reduce absenteeism
10-20% increase in absence is associated with a corresponding CO2 concentration of 1400ppm2
2 minutes to breathe in a lungful of exhaled air from other occupants when indoor CO2 concentration is above 1500 ppm
Improve sleep quality
1000ppm CO2 results in a significant worsening of sleep quality and higher stress levels after a night's rest1
1300ppm CO2 decreases deep sleep duration and causes a significant increase in salivary cortisol1
Enhance cognitive performance
21% decrease in cognitive scores is observed with a 400ppm increase in CO2 concentration3
101% higher cognitive function scores in environments with low VOC, low CO2 and high ventilation rates3
Enabling New Use Cases at Home and at Work
1Kang, Mengyuan, et al. “Ventilation Causing an Average CO2 Concentration of 1,000 ppm Negatively Affects Sleep: A Field-Lab Study on Healthy Young People.” Building and Environment, vol. 249, Elsevier BV, Feb. 2024, pp. 111118–18, https://doi.org/10.1016/j.buildenv.2023.111118. Accessed 25 Mar. 2024.
2Shendell, D. G., Prill, R., Fisk, W. J., Apte, M. G., Blake, D., & Faulkner, D. (2004). Associations between classroom CO2 concentrations and student attendance in Washington and Idaho. Indoor Air, 14(5), 333–341. https://doi.org/10.1111/j.1600-0668.2004.00251.
3Allen, Joseph G., et al. “Associations of Cognitive Function Scores with Carbon Dioxide, Ventilation, and Volatile Organic Compound Exposures in Office Workers: A Controlled Exposure Study of Green and Conventional Office Environments.” Environmental Health Perspectives, vol. 124, no. 6, June 2016, pp. 805–12, https://doi.org/10.1289/ehp.1510037.
Notes: B2

## Slide 19: CO2 is a Major Differentiator
Enabling New Use Cases at Home and at Work
Child & elderly safety
Detect animalities in CO2 concentration patterns, potentially detecting a child or pet left in the car or home
Elderly home pattern tracking to anticipate on animalities, alerting family or caretaker
Taxi virus load control
2 minutes to breathe in a lungful of exhaled air from other occupants when CO2 reaches 1500 ppm
User experience enhancement for customers, making them feel safe when ordering a taxi
Notes: B3

## Slide 20: Residential CO2 Detection
Bedroom: Improve sleep quality by maintaining low CO2 concentrations throughout the night.
Children’s room: Alert parents when air quality decreases in the room of your loved ones.
Dining room: Assist in ventilation optimization based on changes in CO2 levels.
Living room: Rapidly changing temperatures and lower CO2 levels can indicate open windows.
© Copyright Sensirion, Switzerland
20
Use Cases
Notes: C2

## Slide 21: Carbon Dioxide (CO2) & Demand Controlled Ventilation
Energy Efficiency
Between 17 – 23% energy savings via DCV / ERV
Cognitive Performance
-21 % per 400 ppm CO2
CO2
CO2
low
high
400ppm
2000ppm
1000ppm
600ppm
outdoor air
Humans exhale CO2
Human presence
increases CO2 concentration
21
© Copyright Sensirion, Switzerland
Satish, Usha, et al. “Is Co2 an Indoor Pollutant? Direct Effects of Low-To-Moderate CO2 concentrations on Human Decision-Making Performance.” Environmental Health Perspectives, vol. 120, no. 12, Dec. 2012, pp. 1671–77, https://doi.org/10.1289/ehp.1104789.
J. Zhang, G. Liu, R, Lutes, M. Brambley, Energy-saving for occupancy-based control (OBC) of Variable-Air-Volume(VAV) Systems, in PNNL-22072, 2013. pp. 1-79, Virginia, The United States of America.
Notes: B7

## Slide 22: Making The Difference With CO2 Sensing
SCD4x Applications

## Slide 23: Demand-Controlled Ventilation (DCV/VAV)
Selectively reduce fresh air supply based on room usage throughout buildings
Can help to reduce energy consumption of heating, ventilation and air-conditioning (HVAC) systems by 20-50%.1,2,3
In the USA, DCV is mostly regulated by ASHRAE 62.1 in commercial buildings. CO2 measurement accuracy of 30ppm + 3% is required as per 62.1-2022 Addendum d.
Europe is governed by EN 15251, which does not define an accuracy for measurement devices, and local legislation which may include measurement device requirements (e.g. Belgian NBN D50-001).
1Li B and Cai W. A novel CO2-based demand-controlled ventilation strategy to limit the spread of COVID-19 in the indoor environment. Build. Environ. 219 (2022) 109232.
2Merema B, Delwati M, Sourbron M and Breesch H. Demand controlled ventilation (DCV) in school and office buildings: Lessons learnt from case studies. Energy Build. 172 (2018) 349–360.
3Mysen M, Berntsen S, Nafstad P, Schild PG. Occupancy density and benefits of demand-controlled ventilation in Norwegian primary schools. Energy and Buildings 37 (12) (2005), 1234-1240.

## Slide 24: Air Conditioners
Optimize energy consumption
Make your ventilation and cooling energy efficient
Power saving modes can be triggered by CO2-based room presence detection.
Trigger ventilation to turn off when open windows events occur, which can be indicated by changes in temperature and lower CO2 concentrations.
Improve energy efficiency of heat recovery ventilators (HRV) and demand control ventilation (DCV) through smarter regulation loops in residential buildings.
© Copyright Sensirion, Switzerland
24
Notes: C6

## Slide 25: Thermostats
Increase room sensing capabilities beyond temperature and humidity
Expand smart thermostat features through CO2 sensing
CO2 levels can be monitored to drive air conditioning.
Assist in energy efficiency by sensing room inoccupation through low CO2 levels.
Measure CO2 trends in different rooms to establish habits and triggers.
© Copyright Sensirion, Switzerland
25
Notes: C5

## Slide 26: IAQ Monitors & Purifiers
Improve well-being in all indoor environments
CO2 monitoring is a cornerstone of comfortable homes
Sleep quality deteriorates with higher CO2 levels.
Increased remote working practices increase the need for air quality monitoring at home.
Higher CO2 levels can trigger alerts to increase air purification and ventilation.
© Copyright Sensirion, Switzerland
26
Notes: C4

## Slide 27: CO2 Detection for Grain Monitoring
Early Detection of Spoilage detect spoilage much earlier than traditional methods by identifying increased CO2 levels caused by mold or insect
Improved Grain Quality maintain optimal storage conditions, leading to higher quality grain and potentially better market value
Reduced Fumigation Costs Early detection of spoilage allows for timely interventions, reducing the need for extensive and costly fumigation treatments
Sustainable Storage Management Continuous CO2 monitoring helps in managing grain storage more sustainably by preventing spoilage and reducing waste
© Copyright Sensirion, Switzerland
27
Application Benefits
Notes: C9

## Slide 28: ASHRAE 62.1 California Title 24 RESET WELL …
Key Building Standards and CO2 Sensor Requirements

## Slide 29: The four different types of standards…
Standards for Building Codes
→ Legally enforced, regulate how buildings must be built (incl. HVAC / DCV / ERV systems and their components)
…for which CO2 sensors are relevant
Standards for Green Building Certification
→ Voluntary, focus on systems for good building energy-efficiency incl. HVAC/DCV/ERV systems
Standards for Indoor Air Quality Device Certification
→ Voluntary, focus on device-level (sensing) performance (for e.g. IAQ monitors, air purifiers)
Standards for Healthy Building Certification
→ Voluntary, focus on systems for healthy indoor environments incl. indoor air quality monitoring / smart ventilation

## Slide 30: ASHRAE Standards
ANSI/ASHRAE 62.2 Ventilation and Acceptable Indoor Air Quality in Residential Buildings
Defines minimum requirements for natural/mechanical ventilation systems and the building envelope to provide acceptable indoor air quality (IAQ) in residential buildings.
Standard 62.2 in the US is referenced in 19 state building codes.
62.1 Commercial Ventilation & 62.2 Residential Ventilation
The American Society of Heating, Refrigeration and Air-conditioning Engineers (ASHRAE) is a global organization publishing guidelines, funding research projects and offering continuous education to improve building services, engineering and indoor air quality.
ANSI/ASHRAE 62.1 - Ventilation and Acceptable Indoor Air Quality
Specifies ventilation rates and other measures intended to provide acceptable indoor air quality and minimize adverse health effects. The standard also contains requirements related to certain contaminants.
Standard 62.1 in the US is referenced in 18 state building codes (and also in some foreign countries).
https://www.ashrae.org/file library/about/government affairs/advocacy toolkit/virtual packet/standard-62.1-fact-sheet.pdf
https://www.ashrae.org/file%20library/about/government%20affairs/advocacy%20toolkit/virtual%20packet/standard-62.2-fact-sheet.pdf

## Slide 31: ASHRAE 62.1-2022 Addendum d
Context
Officially released March 31st 2025 Correction to Addendum ab from October 2023
Tougher accuracy requirements for CO2 sensors:
  Old: ± 75 ppm up to 1’000 ppm (expanded to 2’500 ppm by Addendum ab)
  New: ± (30 ppm + 3% m.v.) up to 2’500 ppm

## Slide 32: California Title 24
Provides comprehensive building energy efficiency standards for residential and non-residential buildings
“[…] The Standards contain energy and water efficiency requirements (and indoor air quality requirements) for newly constructed buildings, additions to existing buildings, and alterations to existing buildings […]”
Several other US states have adapted (parts) of the Californian Title 24 into their building codes
Section 120.1 – 4 F and Section 160.2 – 5 D:
Apply for CO2-controlled demand-controlled ventilation (DCV):
“CO2 sensors shall be certified by the manufacturer to be accurate within plus or minus 75 ppm at a 600 and 1000 ppm concentration when measured at sea level and 25°C, factory calibrated and certified by the manufacturer to require calibration no more frequently than once every 5 years.”
© Copyright Sensirion, Switzerland
32
Building Energy Efficiency Standards for Residential and Nonresidential Buildings
Notes: E11

## Slide 33: Overview: Building Standards & CO2 Accuracy Requirements
Chart: Accuracy as a Function of CO2 Gas Concentration
| 400 | 500 | 600 | 700 | 800 | 900 | 1000 | 1001 | 1100 | 1200 | 1300 | 1400 | 1500 | 1600 | 1700 | 1800 | 1900 | 2000 | 2001 | 2100 | 2200 | 2300 | 2400 | 2500 | 2600 | 2700 | 2800 | 2900 | 3000 | 3100 | 3200 | 3300 | 3400 | 3500 | 3600 | 3700 | 3800 | 3900 | 4000 | 4100 | 4200 | 4300 | 4400 | 4500 | 4600 | 4700 | 4800 | 4900 | 5000
SCD43 Accuracy Spec | 42 | 45 | 48 | 51 | 54 | 57 | 60 | 60.03 | 63 | 66 | 69 | 72 | 75 | 78 | 81 | 84 | 87 | 90 | 90.03 | 93 | 96 | 99 | 102 | 105 | 108 | 111 | 114 | 117 | 120 | 123 | 126 | 129 | 132 | 135 | 138 | 141 | 144 | 147 | 150 | 153 | 156 | 159 | 162 | 165 | 168 | 171 | 174 | 177 | 180
RESET Grade B | 62 | 65 | 68 | 71 | 74 | 77 | 80 | 80.03 | 83 | 86 | 89 | 92 | 95 | 98 | 101 | 104 | 107 | 110 | 150.05 | 155 | 160 | 165 | 170 | 175 | 180 | 185 | 190 | 195 | 200 | 205 | 210 | 215 | 220 | 225 | 230 | 235 | 240 | 245 | 250 | 255 | 260 | 265 | 270 | 275 | 280 | 285 | 290 | 295 | 300
California Title 24 | 75 | 75 | 75 | 75 | 75 | 75 | 75
WELL | 70 | 75 | 80 | 85 | 90 | 95 | 100 | 100.05 | 105 | 110 | 115 | 120 | 125 | 130 | 135 | 140 | 145 | 150
ASHRAE 62.1 | 42 | 45 | 48 | 51 | 54 | 57 | 60 | 60.03 | 63 | 66 | 69 | 72 | 75 | 78 | 81 | 84 | 87 | 90 | 90.03 | 93 | 96 | 99 | 102 | 105
California Title 24
±75 ppm @ 400 and 1000 ppm CO2
RESET Grade B
±(50 ppm + 3% m.v.) @ 400 - 2000 ppm CO2
±(50 ppm + 5% m.v.) @ 2001 - 5000 ppm CO2
© Copyright Sensirion, Switzerland
33
Chart: CO2 Gas Concentration (ppm)
| 400 | 500 | 600 | 700 | 800 | 900 | 1000 | 1001 | 1100 | 1200 | 1300 | 1400 | 1500 | 1600 | 1700 | 1800 | 1900 | 2000 | 2001 | 2100 | 2200 | 2300 | 2400 | 2500 | 2600 | 2700 | 2800 | 2900 | 3000 | 3100 | 3200 | 3300 | 3400 | 3500 | 3600 | 3700 | 3800 | 3900 | 4000 | 4100 | 4200 | 4300 | 4400 | 4500 | 4600 | 4700 | 4800 | 4900 | 5000
SCD43 Accuracy Spec | 42 | 45 | 48 | 51 | 54 | 57 | 60 | 60.03 | 63 | 66 | 69 | 72 | 75 | 78 | 81 | 84 | 87 | 90 | 90.03 | 93 | 96 | 99 | 102 | 105 | 108 | 111 | 114 | 117 | 120 | 123 | 126 | 129 | 132 | 135 | 138 | 141 | 144 | 147 | 150 | 153 | 156 | 159 | 162 | 165 | 168 | 171 | 174 | 177 | 180
RESET Accuracy Spec | 62 | 65 | 68 | 71 | 74 | 77 | 80 | 80.03 | 83 | 86 | 89 | 92 | 95 | 98 | 101 | 104 | 107 | 110 | 150.05 | 155 | 160 | 165 | 170 | 175 | 180 | 185 | 190 | 195 | 200 | 205 | 210 | 215 | 220 | 225 | 230 | 235 | 240 | 245 | 250 | 255 | 260 | 265 | 270 | 275 | 280 | 285 | 290 | 295 | 300
Old SCD41 Accuracy Spec | 42 | 45 | 48 | 51 | 54 | 57 | 60 | 60.03 | 63 | 66 | 69 | 72 | 75 | 78 | 81 | 84 | 87 | 90 | 90.03 | 93 | 96 | 99 | 102 | 105 | 108 | 111 | 114 | 117 | 120 | 123 | 126 | 129 | 132 | 135 | 138 | 141 | 144 | 147 | 150 | 153 | 156 | 159 | 162 | 165 | 168 | 171 | 174 | 177 | 180
CA T24 Accuracy Spec | 75 | 75 | 75 | 75 | 75 | 75 | 75
WELL Standard | 70 | 75 | 80 | 85 | 90 | 95 | 100 | 100.05 | 105 | 110 | 115 | 120 | 125 | 130 | 135 | 140 | 145 | 150
ASHRAE 62.1 | 42 | 45 | 48 | 51 | 54 | 57 | 60 | 60.03 | 63 | 66 | 69 | 72 | 75 | 78 | 81 | 84 | 87 | 90 | 90.03 | 93 | 96 | 99 | 102 | 105
WELL
±(50 ppm + 5% m.v.) @ 400 - 2000 ppm CO2
ASHRAE 62.1-2022 Addendum d*
±(30 ppm + 3% m.v.) @ 600, 1000 and 2500 ppm CO2
*ASHRAE 62.1-2022 excl. addenda (and previous) set the same CO2 sensor accuracy requirements as California Title 24. U.S. Building Codes generally take several years to adopt updates to referenced standards. USGBC LEED and BREEAM certification systems also reference ASHRAE Standard 62.1
Notes: E9

## Slide 34: SCD4x Sensor Evaluation Kit
Get ready for take-off with SCD4x

## Slide 35: How to get started: SEK-SCD4x-Sensor (Evaluation Kit)
The SEK-SCD41 evaluation kit represents the perfect toolkit for effective SCD4x sensor evaluation and efficient prototyping.”
“
© Copyright Sensirion, Switzerland
35

## Slide 36: Connecting your SCD4x Development Board to a computer
Sensirion Sensor Bridge
(not included in SEK-SCD4x)
Computer
Windows/Mac/Linux
USB port
Adapter Cable
(included in SEK-SCD4x)
SCD4x Development Board
(included in SEK-SCD4x)
© Copyright Sensirion, Switzerland
36

## Slide 37: SCD4x CO2 Sensor Evaluation Kit: ControlCenter Tutorial
 Click here to watch tutorial (https://www.youtube.com/watch?v=siraFIX7JHk)
© Copyright Sensirion, Switzerland
37

## Slide 38: Resources: www.sensirion.com/my-scd-ek
Quick Start Guide Evaluation Kit
Documentation
Key Documents Evaluation Kit SEK-SCD41
SEK-SCD41 Technical Description
SEK-SCD41 STEP File
Key Documents for SCD4x
Datasheet SCD4x
SCD4x STEP file
SCD4x Altium Footprint
Handling Instructions
Design-In Guide
Testing Guide
Overview: The most important Design-in Recommendations
Good coupling to ambient air
Isolation from air turbulences
a
b
Decoupling from vibration sources
Decoupling from external heat sources
c
d
Shielding from direct sunlight
Ensure stable supply voltage
e
f
© Copyright Sensirion, Switzerland
38

## Slide 39: Resources: www.sensirion.com/my-scd-ek
SCD41 Development Board Interfacing
SCD4x Embedded Driver
SCD4x Arduino Driver
SCD4x RaspberryPi Driver
Arduino Examples
SCD4x Makecode for micro:bit
SEK-ControlCenter Viewer Software
Drivers
Quick Start Guide Evaluation Kit
Documentation
© Copyright Sensirion, Switzerland
39

## Slide 40: Resources: www.sensirion.com/my-scd-ek
Quick Start Guide Evaluation Kit
Sensor settings
RH Output
T Output
CO2 Output
© Copyright Sensirion, Switzerland
40

## Slide 41: Resources: www.sensirion.com/my-scd-ek
SCD41 Development Board Interfacing
Arduino
RaspberryPi
Python Package
Use Case Tutorials
BLE Gadget for CO2 Monitoring
Drivers
Quick Start Guide Evaluation Kit
Documentation
Tutorials
© Copyright Sensirion, Switzerland
41

## Slide 42: SCD4x CO2 Sensor Family
Take-home Message

## Slide 43: Five reasons why the SCD4x is ahead of its competitors
Highly selective photoacoustic measurement  Accurate and reliable sensor output
SMD solderable, tape & reel package  Highly automated manufacturing
Completely encapsulated package  Robust against mechanical stress & dust
Smallest optical CO2 sensor  Maximum freedom of design
Large supply voltage range: 2.4 V – 5.5 V
 Easy integration & Design-in
© Copyright Sensirion, Switzerland
43